## Slide 1
Appraisal of real-time bioaerosol detection techniques and classification algorithms
Ian Crawford, David Topping, Martin Gallagher
Simon Ruske & Virginia Foot
International Aerosol Modeling Algorithms Conference 2019
UC Davis
0000-0003-4433-7310
i.crawford@manchester.ac.uk

## Slide 2
Primary biological aerosol particles (PBAP) are ubiquitous in the atmosphere and have significant impacts on plant, animal and human health. They also may play an important role in cloud-aerosol interactions through ice nucleation processes, thus quantifying PBAP emissions is critical for understanding their impact in a changing climate. Despite this need, accurately classifying and quantifying atmospheric PBAP remains a significant technical challenge, limiting their inclusion in climate and health models.
Emerging commercial real-time ultraviolet light induced fluorescence instrumentation offers high spatiotemporal resolution compared to traditional offline bioaerosol detection techniques, however, PBAP classification with these methods requires a careful and considered approach to data processing to produce high quality outputs.
Here we present and evaluate a variety of different data processing approaches, including supervised and unsupervised machine learning methodologies. We appraise the current state of technological outputs and provide potential end-users with a road map of future development, focusing on data product quality.
Abstract
Notes: BULLET POINTS!

## Slide 3
UV-LIF
Pohlker et al. 2012

## Slide 4
Spectrometers
WIBS
3 channels
Excitation 280/370 nm
2 detection bands
310-400/420-650nm
Sizing: 0.8 – 20 µm
Shape: quadrant detector
MBS
8 channels
Excitation 280nm
8 detection bands
310 – 638 nm
Sizing: 0.8 – 20 µm
Shape: dual CMOS array
PLAIR Rapid-E
32 channels
Excitation: 337 nm
32 detection bands: 350-800 nm
4 FL ifetime decay bands
Sizing: 0.5-100 µm
Shape: time resolved scattering
DMT SIBS
28 channels
Excitation 285/370 nm
16 detection bands: 300-720 nm
Sizing: 0.5-30 µm
Shape: quadrant detector
Increasing Complexity & Specificity

## Slide 5
Könemann et al. 2019

## Slide 6
Historic Analysis Methods
Savage et al., 2013
Hernandez et a., 2016

## Slide 7
Current Analysis Methods
Machine
Learning
Unsupervised
Clustering
Dimensionality
Reduction
Supervised
Classification
Regression
Hierarchical
Agglomerative
Clustering
K-Means
Density Based
Spectral Clustering
and Noise
Gradient
Boosting
Decision
trees
Gaussian
Mixture
K-nearest
neighbour
SVM
Artificial
Neural
networks
Fuzzy C-Means
SLOW
FAST
Ruske et al., 2018

## Slide 8
Hierarchical Agglomerative Clustering
No a priori information needed but highly sensitive to pre-processing
Need to choose best solution – e.g., Calinski–Harabasz index
Groups similar data points into clusters

## Slide 9
Huffman et al., 2013

## Slide 10
Density Based Spectral Clustering and Noise
Groups similar data points into clusters parametrically
Define neighbourhood radius and minimum radial population
Intrinsically calculated optimal solution
Particles may be left unclassified
New promising method – requires validation for ambient sampling

## Slide 11
Decision tree – Gradient Boosting
Fit one tree to data
Evaluate where performing well
Fit second tree where model performs poorly
Repeat and add trees
MBS - London

## Slide 12
Algorithm summary

## Slide 13
Challenges - Calibration
Robinson et al., 2017
Need suitable fluorescent
calibration standards
Need calibration methods
for multichannel detectors
Könemann et al. 2019

## Slide 14
Challenges – Reference Data
Need a much larger reference database for training
Effect of processing and mixing uncharacterised
Bacteria
Fungal
Pollen

## Slide 15
Challenges – Improving Spectrometers
MBS CMOS
RAPID-E